## Moisture Analyser
Moisture Analyser Restoration Projects Dashboard
Location | Tag No | Equipment Detail |  | Current Problem | Current Status | Recommendation | Action Plan | Responsible Party | Start | Finish | Remark
|  | S/n | SW rev
Agbada | A-26QZA-351 | 1167 | T.cod | TECSPT at limit fault | Unit is still in service at customer site but NOT in good functional state | Send unit back to factory for repair/upgrade and recalibration | To send unit back to GE factory | ME | 30/4/2017 | 30/11/2017 | Unit sent to factory. PO: 4510382079. OEM requested more time to deliver hence the date was extended to 30th Nov'17
[7 empty] | Upgrade and recalibrate unit and return back to SPDC/GCA | ME
[7 empty] | Install repaired unit on Site | ME
[7 empty] | Commission the unit on site & drop pressure at sample take off point to 3-5 | ME
Imor River | I-26QZA-351 | 1168 | K.cod | Detached mirror due to prolonged exposure to condensate | Unit is NOT in service but in SPDC possession after de-commissioning | Send unit back to factory for upgrade and recalibration | To send unit back to GE factory | ME | 30/4/2017 | 30/11/2017 | Unit sent to factory. PO: 4510382079. OEM requested more time to deliver hence the date was extended to 30th Nov'17
Upgrade and recalibrate unit and return back to SPDC/GCA
|  | 1379 | T.cod | Eroded mirror surface due to exposure to liquid carry-over | Unit is still in service at customer site but NOT in good functional state | Send unit back to factory for repair and recalibration | To send unit back to GE factory | ME | 30/4/2017 | 30/11/2017 | Unit sent to factory. PO: 4510382079. OEM requested more time to deliver hence the date was extended to 30th Nov'17
Repair and recalibrate unit and return back to SPDC/GCA
[7 empty] | Install repaired unit on Site | ME |  |  | Unit sent to factory. PO: 4510382079. OEM requested more time to deliver hence the date was extended to 30th Nov'17
Commission the unit on site
[6 empty] | Install Pre-conditioning Sample system to prevent issues with liquid carry-over | Place order for Pre-conn SS | ME
Deliver Pre-conn SS to customer
[7 empty] | Install and commission Pre-con sample system | ME |  |  | PO: 4510382079
Obigbo AGG | 26-QIT-093 | 1119 | T.cod | Meter in fault mode (TEC SET POINT OUT OF RANGE). | Meter is running on outdated software which could be reason behind it's signal loss | OEM should be contacted to download latest software on the meter | Bring OEM for further troubleshooting | ME | 25/10/2017 | 27/10/2017 | Vendor mobilised 1st Nov'17. 48hrs obersavation test completed.
Obigbo NAG | NA |  | M.cod | Meter not in service. 5Amp fuse blown. Power card is short circuit. Dew point values vary (increase significantly) in the afternoon irrespective of weather condition | Power supply board damaged due to power upset at Obigbo | Replace power supply board | Procure power supply board WO 22588946 Provide scan display sketch alongside the other values as requested by the OEM for further analysis of the varying dew point values. | ME | 19/9/2017 | 2017-05-12 | PO 4510389177 issued to the vendor. Awaiting OEM to confirm the deliver date
Obigbo Sales Manifold |  |  | T.cod | Meter not in service. 5Amp fuse blown. Power card is short circuit | Not in service due to power card damage. 240VDC supply available to the meter | Propose to procure 2 new units of Aurora moisture analyser | PO to be advised by mtce Exec 18/11/2017 | ME | 19/9/2017 | 20/12/2017 | WO 22588946.